# Emergency Response Plan

Contents

Purpose	1

Scope	2

Responsibilities	2

Supervisors	2

Workers	2

Safety Offices	2

Emergency Guidebook	3

Emergency Codes for Hospital-Hosted Locations	3

Accidents and Incidents	4

Post Emergency Inventory Check	4

Medical Emergencies	4

Fire	4

Chemical Spills	5

Biological Spills	5

Power Failure	5

Failure of Primary Containment Devices	5

Loss of Containment	6

Emergency Egress	6

Notification of Key Personnel	6

Natural Disasters and Human Caused Disasters	6

Incident Follow-up and Recommendations to Mitigate Future Risks	6

Training	6

# Purpose

<<<< This SOP is left in DOCX format so that you may edit it for your own laboratory. YOU MUST EDIT THIS SOP>>>

The purpose of this SOP is to lay out the responsibilities and procedures required for emergency response.

# Scope

This SOP applies to all laboratory occupants and their Supervisors.

# Responsibilities

## Supervisors

Supervisors are responsible for:

- Reviewing this SOP on a regular basis. Review is to consider and mitigate the risks of spill, loss of containment and exposure or other harm. Refer to Performing Risk Assessments SOP.

- Ensuring that all workers under their supervision are trained on and are proficient in performing the steps of this SOP.

## Workers

Workers are responsible for:

- Following this SOP as approved by their supervisor

# Safety Offices

There are four safety offices for McMaster University:

- Environmental and Occupational Health Support Services, EOHSS

  - Serves all faculties except the Faculty of Health Sciences

  - Occupational health and safety, excluding radioisotopes and biohazards

  - eohss@mcmaster.ca

  - http://www.workingatmcmaster.ca/eohss

- Faculty of Health Sciences Safety Office

  - Serves only the Faculty of Health Sciences

  - Occupational health and safety, excluding radioisotopes and biohazards

  - fhsso@mcmaster.ca

  - http://fhs.mcmaster.ca/safetyoffice

- Health Physics Department

  - Serves all faculties

  - Management of radioisotopes

  - hphys@mcmaster.ca

  - http://www.mcmaster.ca/healthphysics/index.html

- McMaster Biosafety Office

  - Serves all faculties

  - Management of biohazards

  - robertjv@mcmaster.ca

  - https://biosafety.mcmaster.ca

# Emergency Guidebook

- All laboratory occupants and all Supervisors must read and familiarize themselves with the Emergency Guidebook, particularly emergency contact numbers and locations.

- http://security.mcmaster.ca/campus_emergencies_guide.html

- If there is information that you require, that is not in the emergency guidebook, please contact your Departmental Chair or Area Manager for information.

- Add required information to this SOP.

# Emergency Codes for Hospital-Hosted Locations

# Accidents and Incidents

- All accidents and incidents must follow this same general procedure.

- Control the situation, prevent escalation of danger.

- Inform nearby occupants and Supervisor.

- Gather required information for a medical visit (MSDS or PSDS or similar).

- Seek medical attention if necessary.

- http://www.workingatmcmaster.ca/med/document/RMM-1000-Reporting-and-Investigating-Injury-Incident-Occupational-Disease-Program-1-36.pdf

- Fill an incident report. Signed by Supervisor. Signed by Dept Chair.

- Submit incident report to EOHSS (non-FHS) or FHS Safey Office (FHS).

# Post Emergency Inventory Check

- When any emergency causes evacuation, upon re-entry to the building, all storage units of pathogens should be checked for integrity and theft of inventory items.

- Any damage or theft or suspicion of theft should be reported immediately to the Supervisor and to the Biosafety Office robertjv@mcmaster.ca x23453

# Medical Emergencies

- See Emergency Guidebook for location-specific information.

- http://security.mcmaster.ca/campus_emergencies_guide.html

# Fire

- All laboratory occupants are to participate in Fire Safety training on Mosaic, which is appropriate to their work environment.

- http://mosaic.mcmaster.ca

- Fire safety procedures are DIFFERENT in hospitals versus non-hospital hosted locations. Some may require non-hospital hosted training, others may require hospital-hosted training and some may require both.

- Annual updates may be required.

- Collect and update training records annually.

# Chemical Spills

- All persons using chemicals are to participate in Chemical Handling and Spills training through Mosaic.

- http://mosaic.mcmaster.ca

- https://fhs.mcmaster.ca/safetyoffice/documents/UniversalPrecautionsandChemicalSpillpage.pdf

- http://cll.mcmaster.ca/eohss/chemhandlingspills.htm

- http://www.workingatmcmaster.ca/med/document/Lab-Safety-Handbook-1-36.pdf

- Collect and update training records annually.

# Biological Spills

- All persons using biohazards must participate in Biosafety training through Mosaic and obtain related, practical training within their laboratory.

- http://mosaic.mcmaster.ca

- Biosafety spill protocols should be contained within their laboratory’s biosafety manual and should be signed off by all lab occupants.

# Power Failure

- In the event of a power failure that affects primary containment devices follow the procedure below.

- In the event of a power failure, emergency power is supplied.

- Stop work, stand by and wait for instructions from your Supervisor.

# Failure of Primary Containment Devices

- Failure of a BSC would be demonstrated by loss of power, alarming, sudden change in magnehelic readings or any sound or vibration that would indicate a malfunction.

- The following steps are the same if the BSC can remain on or if it must be turned off. They key is to close all containers as soon as possible.

- Stop all work.

- Close all sample containers.

- Spray out all sample containers.

- Decontaminate all surfaces.

- Collect all waste into waste bag.

- Change gloves.

- Close waste bag and spray all surfaces.

- Spray gloves and leave them in BSC.

- Wait minimum contact time for decontamination solution.

- Turn off BSC.

- Discard gloves in regular waste.

- Transport bagged waste to biohazard waste box or bin.

- Wash hands.

- Sign the BSC as out of order.

- Report the malfunction immediately to your Supervisor.

# Loss of Containment

- Loss of containment constitutes a spill and should be dealt with according to the type of spill encountered. Spill SOPs are found where in your laboratory? SOP #?

- Loss of containment may also constitute an exposure and should be dealt with according to exposure protocols. Exposure protocols are found were in your laboratory? SOP#?

# Emergency Egress

- Emergency egress is the path you take to evacuate your laboratory.

- Where is your laboratory’s ‘meeting place’?

- How do you get to your meeting place?

# Notification of Key Personnel

- Personnel to be notified in emergencies should be included in each emergency SOP.

- Supervisor’s name, extension and emergency #

# Natural Disasters and Human Caused Disasters

- These disasters are captured in the Crisis Management Response Plan.

- http://www.workingatmcmaster.ca/med/document/RMM-1200-Crisis-Management-Response-Plan-1-36.pdf

# Incident Follow-up and Recommendations to Mitigate Future Risks

- All incidents are to be reported according to RMM1000. Ensure a copy of RMM1000 is printed and in your laboratory biosafety manual.

- http://www.workingatmcmaster.ca/med/document/RMM-1000-Reporting-and-Investigating-Injury-Incident-Occupational-Disease-Program-1-36.pdf

- Accident followup instructions are found in RMM1000 and include:

  - Scene assessment

  - Interviewing

  - Identify contributing factors

  - Make recommendations for corrective actions

  - Ensure recommendations are completed

# Training

- Ensure all lab members have read and understood this SOP.

- Record dates and signatures of training.

- Follow your lab protocol for in-house training i.e. what signatures are required.

- If this SOP is updated, all persons require to be re-trained and re-sign.